## Slide 1: Technical Challenges in SMBG
Sara Sedaghat M.D.
Gabric Diabetes Education Association
Medical Affaires Department, Farir Asa Teb

## Slide 2: Case 1:
A 10 year old boy diagnosed with T1DM since he was 6 years old.
Weight: 35 kg
Height: 143 Cm
Tx Plan:
  Lantus 16 unit- 6:00 A.M.
  Apidra: 4 unit pre-meal
His recent HbA1c: 9.8%

## Slide 3: HbA1c & Average Glucose (AG)
HbA1c: 9.8 %
AG:212-240 mg/dL

## Slide 4: SMBG log book data
Blood Glucose Average: 155 mg/dL

## Slide 5: SMBG and HbA1c Mismatch
SMBG Frequency and Timing
Shared glucometer and data

## Slide 6: His 7-Point SMBG report has been shown:
صبحانه
نهار
شام
قبل از خواب
صبحانه
نهار
شام
قبل از خواب
230
130
85
220
120
188
107
صبحانه
نهار
شام
قبل از خواب
صبحانه
نهار
شام
قبل از خواب
212
152
80
314
130
132
157
صبحانه
نهار
شام
قبل از خواب
صبحانه
نهار
شام
قبل از خواب
165
140
81
290
80
172
132
روز اول
روز سوم
روز دوم
Notes: 6

## Slide 7: SMBG and HbA1c Mismatch
Inaccurate HbA1c
Inaccurate SMBG result
Patient’s specific conditions:
    Hemolytic diseases
    Recent large blood loss
    Renal failure

## Slide 8: His 7-Point SMBG report with new glucometer:
صبحانه
نهار
شام
قبل از خواب
صبحانه
نهار
شام
قبل از خواب
230
220
50
300
180
388
250
صبحانه
نهار
شام
قبل از خواب
صبحانه
نهار
شام
قبل از خواب
60
270
200
340
210
400
260
صبحانه
نهار
شام
قبل از خواب
صبحانه
نهار
شام
قبل از خواب
280
250
189
290
220
272
237
روز اول
روز سوم
روز دوم
Notes: 8

## Slide 9: SMBG log book data
Blood Glucose Average: 242 mg/dL

## Slide 10: What’s your judgement and explanation?
Misleading Data

## Slide 11: Misleading SMBG Data
Poor Performance of Blood Glucose Meters

## Slide 12: Poor Performance of Blood Glucose Meters
Higher HbA1c levels
Increased hypoglycemia incidence
Insulin dosing error
Increased Glycemic Variability (GV)
Failure to detect hypoglycemia
Adverse Clinical & Economical Outcomes:

## Slide 13: American Consensus on Accuracy of SMBG
Up to 50% of SMBG readings vary from true value more than 20%.
ADA consensus panel:
American Diabetes Association. Self-Monitoring of Blood Glucose. Diabetes Care. 1990;13(Supplement 1):41-6.
Notes: ADA consensus panel گزارش کرد که نیمی از SMBG هایی که در امریکا و اروپا انجام می شود، بیشتر از 20% خطا دارد که در ایران این ارقام بسیار بالاتر است.

## Slide 14
Operational factors:
Human Errors
Meter-Inherent Factors:
BGM Accuracy
(Erbach, Freckmann et al. 2016)
Notes: Inappropriate handling of SMBG has been identified as the most common factor affecting BG results; more than 90% of overall inaccuracies result from incorrect use of BG meters.
Ref:
Erbach, M., G. Freckmann, R. Hinzmann, B. Kulzer, R. Ziegler, L. Heinemann and O. Schnell (2016). "Interferences and Limitations in Blood Glucose Self-Testing: An Overview of the Current Knowledge." J Diabetes Sci Technol 10(5): 1161-1168.
Approximately 91–97% of overall inaccuracies are operator dependent.
Ref:
Tonyushkina, K. and J. H. Nichols (2009). "Glucose meters: a review of technical challenges to obtaining accurate results." J Diabetes Sci Technol 3(4): 971-980.

## Slide 15: Meter-Inherent factors

## Slide 16: Analytical Performance: Accuracy Criteria
The standard ISO 15197:2015 has been mandatory in the EU countries since June 2016.
International Organization for Standardization. In vitro diagnostic test systems—requirements for blood-glucose monitoring systems for self-testing in managing diabetes mellitus. EN ISO/DIS 15197:2013. Available at: http:// www.iso.org/iso/home/store/catalogue_tc/catalogue_detail.htm?csnumber=54976.

## Slide 17: New ISO: ISO 15197:2003 VS ISO 15197:2013

## Slide 18: Clark Error Grid (CEG) Zones

## Slide 19: Reduction in BGM Error from 20% to 15%
Schnell, O., & Erbach, M. (2014). Impact of a reduced error range of SMBG in insulin-treated patients in Germany. Journal of diabetes science and technology, 8(3), 479-482.
Notes: Reduction in Blood Glucose Monitoring Error from 20% (ISO 15197:2003) to 15% (ISO 15197:2015)

## Slide 20
Operational factors:
Human Errors
Meter-Inherent Factors:
BGM Accuracy
Erbach, M., Freckmann, G, et al (2016). Interferences and Limitations in Blood Glucose Self-Testing: An Overview of the Current Knowledge. J Diabetes Sci Technol, 10(5), 1161-1168.
Notes: Inappropriate handling of SMBG has been identified as the most common factor affecting BG results; more than 90% of overall inaccuracies result from incorrect use of BG meters.
Ref:
Erbach, M., G. Freckmann, R. Hinzmann, B. Kulzer, R. Ziegler, L. Heinemann and O. Schnell (2016). "Interferences and Limitations in Blood Glucose Self-Testing: An Overview of the Current Knowledge." J Diabetes Sci Technol 10(5): 1161-1168.
Approximately 91–97% of overall inaccuracies are operator dependent.
Ref:
Tonyushkina, K. and J. H. Nichols (2009). "Glucose meters: a review of technical challenges to obtaining accurate results." J Diabetes Sci Technol 3(4): 971-980.

## Slide 21: Operational Factors

## Slide 22: Human Factors With a Potential Impact on the Analytical Performance of SMBG
Incorrect use of BG meters
Incorrect performance of coding
Inappropriate storage and usage of test strips
Inappropriate education of patients and diabetes team
Erbach, M., Freckmann, G, et al (2016). Interferences and Limitations in Blood Glucose Self-Testing: An Overview of the Current Knowledge. J Diabetes Sci Technol, 10(5), 1161-1168.
Notes: Factors With a Potential Impact on the Analytical
Performance of SMBG.
Human:
Incorrect use of BG meters
Incorrect performance of coding
Inappropriate storage and usage of test
strips
Inappropriate education of patients and
diabetes team
Ref:
Erbach, M., G. Freckmann, R. Hinzmann, B. Kulzer, R. Ziegler, L. Heinemann and O. Schnell (2016). "Interferences and Limitations in Blood Glucose Self-Testing: An Overview of the Current Knowledge." J Diabetes Sci Technol 10(5): 1161-1168.

## Slide 23: Human Factors With a Potential Impact on the Analytical Performance of SMBG
Incorrect use of BG meters
Incorrect performance of coding
Inappropriate storage and usage of test strips
Inappropriate education of patients and diabetes team
Erbach, M., Freckmann, G, et al (2016). Interferences and Limitations in Blood Glucose Self-Testing: An Overview of the Current Knowledge. J Diabetes Sci Technol, 10(5), 1161-1168.
Notes: Factors With a Potential Impact on the Analytical
Performance of SMBG.
Human:
Incorrect use of BG meters
Incorrect performance of coding
Inappropriate storage and usage of test
strips
Inappropriate education of patients and
diabetes team
Ref:
Erbach, M., G. Freckmann, R. Hinzmann, B. Kulzer, R. Ziegler, L. Heinemann and O. Schnell (2016). "Interferences and Limitations in Blood Glucose Self-Testing: An Overview of the Current Knowledge." J Diabetes Sci Technol 10(5): 1161-1168.

## Slide 24: Incorrect use of BG meters
Ferretti, J., & Martin, K. (2008). Isopropyl alcohol left on the skin falsely lowers capillary glucose values. JCOM-WAYNE PA-, 15(4), 179.

## Slide 25: External pressure may lead to unreliable readings
Different external pressures led to ≥10% differences in glucose concentrations in in 5%–13% of the participants.
On average, blood sugar readings were lower when people put pressure on the finger.
Hortensius, J., Slingerland, R. J.,et al(2011). Self-Monitoring of Blood Glucose: The Use of the First or the Second Drop of Blood. Diabetes care, 34(3), 556-560.
Notes: External pressure may lead to unreliable readings.
On average, blood sugar readings were lower when people put pressure on the finger.
Different external pressures led to ≥10% differences in glucose concentrations in 5%–13% of the participants.
Ref:
Hortensius, J., R. J. Slingerland, N. Kleefstra, S. J. J. Logtenberg, K. H. Groenier, S. T. Houweling and H. J. G. Bilo (2011). "Self-Monitoring of Blood Glucose: The Use of the First or the Second Drop of Blood." Diabetes Care 34(3): 556-560.

## Slide 26: Hand washing has always been a problem!
With a sample of 0.3 μl, 1 μg of glucose (the weight of a dust particle) will increase the blood glucose by 300 mg/dl.
Ginsberg, B. H. (2009). Factors Affecting Blood Glucose Monitoring: Sources of Errors in Measurement. J Diabetes Sci Technol, 3(4), 903-913.
Notes: Hand washing has always been a problem, but now, with microsample meters, the smallest amount of contaminant can significantly raise the blood glucose. With a sample of 0.3 μl, 1 μg of glucose (the weight of a dust particle) will increase the blood glucose by 300 mg/dl.
Ref:
Ginsberg, B. H. (2009). "Factors Affecting Blood Glucose Monitoring: Sources of Errors in Measurement." J Diabetes Sci Technol 3(4): 903-913.

## Slide 27: Alternate Site Testing (AST)
Notes: برخی دستگاه ها این امکان را دارند که از جاهای دیگر غیر از نوک انگشت (AST) هم خون بگیرند. فقط اشکالش این است که مانند نوک انگشت، قند لحظه را گزارش نمی کند.
Blood Flow
»Fingers = 33 ±10 mL / g •min at 19-22 °C
»Forearm = 4-6 mL / g •min at 19-22 °C
»Abdomen = 4-6 mL / g •min at 19-22 °C
Relative nerve tissue density
»Fingers: Highly enervated (pain issues)
»Forearm: 20-25% of finger nerve levels
»Palm: 75% of finger nerve density
Donald R. Parker, PhD, DABCC, FACB,Chief Scientific Office,Bayer Diabetes Care

## Slide 28: Risk Delay of AST
Jungheim, K., & Koschinsky, T. (2001). Risky delay of hypoglycemia detection by glucose monitoring at the arm. Diabetes care, 24(7), 1303-1304.
Notes: برخی دستگاه ها این امکان را دارند که از جاهای دیگر غیر از نوک انگشت (AST) هم خون بگیرند. فقط اشکالش این است که مانند نوک انگشت، قند لحظه را گزارش نمی کند.
Blood Flow
»Fingers = 33 ±10 mL / g •min at 19-22 °C
»Forearm = 4-6 mL / g •min at 19-22 °C
»Abdomen = 4-6 mL / g •min at 19-22 °C
Relative nerve tissue density
»Fingers: Highly enervated (pain issues)
»Forearm: 20-25% of finger nerve levels
»Palm: 75% of finger nerve density
Donald R. Parker, PhD, DABCC, FACB,Chief Scientific Office,Bayer Diabetes Care

## Slide 29: When it's NOT okay to check from AST
Hypoglycemia symptoms
Hypoglycemia unawareness
Critical Tasks (e.g. Driving)
When blood glucose values are changing rapidly (e.g. exercise, after meal)
If the results of a blood glucose test don't match the way you feel
Notes: Some blood glucose meters allow you to use a blood sample from a nonfingertip or alternate site such as your palm, forearm or upper arm.2 Because the skin of these areas contains fewer nerves than the fingertip, alternate site testing may be more comfortable.3
It's important to know that while blood from your fingertip can be tested at any time, there are times when alternate site testing may not give you the most accurate result.
When it's okay—and not okay—to check from alternate sites
Alternate site testing should only be used when your blood sugar is stable. For example:4
Immediately before a meal
When fasting
Near bedtime
Always check from your fingertip, however, when blood sugar may be changing:4
Following a meal, when blood sugar is rising quickly
After exercise
Whenever you think your blood sugar might be low or falling
Important alternate site testing notes
If you're considering alternate site testing to check your blood glucose, please remember:
Never ignore the symptoms of low or high blood sugar.
If the results of a blood glucose test don't match the way you feel, confirm with a fingertip test. If the fingertip result still doesn't seem to reflect the way you feel, get in touch with your healthcare professional.
Please talk to your healthcare professional before using sites other than your fingertip for testing blood sugar.
1Ranked first most often versus leading competitors. Data on file.
2Talk with your healthcare professional before deciding if alternate site testing is right for you.
3New Scientist. Fingertips and forehead are most sensitive to pain. Available at: https://www.newscientist.com/article/dn25688-fingertips-and-forehead-are.... Accessed March 14, 2016.
4Bina DM, et. al. Clinical impact of prandial state, exercise, and site preparation on the equivalence of alternative-site blood glucose testing. Diabetes Care. 2003; 26:981-985. Available at: http://care.diabetesjournals.org/content/26/4/981.full. Accessed March 14, 2016.

## Slide 30: Human Factors With a Potential Impact on the Analytical Performance of SMBG
Incorrect use of BG meters
Incorrect performance of coding
Inappropriate storage and usage of test strips
Inappropriate education of patients and diabetes team
Erbach, M., Freckmann, G, et al (2016). Interferences and Limitations in Blood Glucose Self-Testing: An Overview of the Current Knowledge. J Diabetes Sci Technol, 10(5), 1161-1168.
Notes: Factors With a Potential Impact on the Analytical
Performance of SMBG.
Human:
Incorrect use of BG meters
Incorrect performance of coding
Inappropriate storage and usage of test
strips
Inappropriate education of patients and
diabetes team
Ref:
Erbach, M., G. Freckmann, R. Hinzmann, B. Kulzer, R. Ziegler, L. Heinemann and O. Schnell (2016). "Interferences and Limitations in Blood Glucose Self-Testing: An Overview of the Current Knowledge." J Diabetes Sci Technol 10(5): 1161-1168.

## Slide 31: The Coding Procedure: Potential Source of Error
Coding is needed to transfer information from the test strip calibration to the BG meter.
Incorrect coding may lead to measurement errors of ±30% or more
Baum, J. M., Monhaut, N et al. (2006). Improving the quality of self-monitoring blood glucose measurement: a study in reducing calibration errors. Diabetes technology & therapeutics, 8(3), 347-357.
Notes: The coding procedure is another potential source of error. Coding is needed to transfer information from the test strip
calibration to the BG meter. Incorrect coding may lead to measurement errors of ±30% or more.21 However, most
modern BG systems no longer require a coding step.

## Slide 32: Miscoding Happens Regularly
Raine CH. Self-monitored blood glucose: A common pitfall. Endo Pract. 2003;9:137-9.
Kristensen G, Nerhus K, Thue G, Sandberg S. Standardized evaluation of instruments for self-monitoring of blood glucose by patients and a technologist. Clin Chem. 2004;50:1068-1071.
Notes: در مطالعه ای که در آمریکا و اروپا انجام شد مشخص شد که 16% از بیماران آموزش دیده دچار خطای کدگذاری می شوند (یا اشتباه انجام می دهند یا فراموش می کنند).

## Slide 33: Effect of miscoding on accuracy
Charles H. Raine, I., et al (2007). Significant Insulin Dose Errors May Occur if Blood Glucose Results are Obtained from Miscoded Meters. Journal of diabetes science and technology, 1(2), 205-210.

## Slide 34: Miscoding Causes Inaccurate Results
Baum JM, Monhaut NM, Parker DR, Price CP. Improving the quality of self-monitoring blood glucose measurement: A study in reducing
calibration errors. Diabetes Technology and Therapeutics. 2006;8:237-247.

## Slide 35: Miscoding Causes Insulin Dose Errors
Charles H. Raine, I., Schrock, L. E., et al. (2007). Significant Insulin Dose Errors May Occur if Blood Glucose Results are Obtained from Miscoded Meters. Journal of diabetes science and technology, 1(2), 205-210.
Notes: این امر سبب می شود که در 50% موارد 2± واحد در تزریق انسولین اشتباه شود. که در مورد دستگاه های اتوکدینگ این خطای دوز نامناسب انسولین بسیار ناچیز است (1.4%)

## Slide 36: Human Factors With a Potential Impact on the Analytical Performance of SMBG
Incorrect use of BG meters
Incorrect performance of coding
Inappropriate storage and usage of test strips
Inappropriate education of patients and diabetes team
Erbach, M., Freckmann, G, et al (2016). Interferences and Limitations in Blood Glucose Self-Testing: An Overview of the Current Knowledge. J Diabetes Sci Technol, 10(5), 1161-1168.
Notes: Factors With a Potential Impact on the Analytical
Performance of SMBG.
Human:
Incorrect use of BG meters
Incorrect performance of coding
Inappropriate storage and usage of test
strips
Inappropriate education of patients and
diabetes team
Ref:
Erbach, M., G. Freckmann, R. Hinzmann, B. Kulzer, R. Ziegler, L. Heinemann and O. Schnell (2016). "Interferences and Limitations in Blood Glucose Self-Testing: An Overview of the Current Knowledge." J Diabetes Sci Technol 10(5): 1161-1168.

## Slide 37: Inappropriate storage and usage of test strips
Inadequate blood volume
Temperature
Test strip reuse
Expiry date
Humidity
Mechanical stress
Chemical stress (reagent)
Original packaging
Erbach, M., Freckmann, G, et al (2016). Interferences and Limitations in Blood Glucose Self-Testing: An Overview of the Current Knowledge. J Diabetes Sci Technol, 10(5), 1161-1168.
Notes: The impact of inappropriate storage of test strips on BG readings is demonstrated by the case of a72-year-old Japanese patient with type 2 diabetes. To simplify the SMBG procedure, he removed the test strips from the packaging and stored them, together with the BG meter, in a small pouch. Due to repeated pseudohypoglycemic readings, the patient abandoned his antihyperglycemic medication, thereby reaching a BG level of 21.8 mmol/L (393 mg/ dL)
Ref: Erbach, M., G. Freckmann, R. Hinzmann, B. Kulzer, R. Ziegler, L. Heinemann and O. Schnell (2016). "Interferences and Limitations in Blood Glucose Self-Testing: An Overview of the Current Knowledge." J Diabetes Sci Technol 10(5): 1161-1168.
Kirk, J. K. and J. Stegner (2010). "Self-Monitoring of Blood Glucose: Practical Aspects." Journal of Diabetes Science and Technology 4(2): 435-439.

## Slide 38: Inappropriate storage and usage of test strips
Issue:
20% of people have difficulties in filling test strips.
Challenge:
Inaccurate result
Pain
solution:
Blood Volume Control (BVC)
Minimal blood sample
Second chance Sampling
Inadequate blood volume
William A Alto et al, Assuring the Accuracy of Home Glucose Monitoring, Diabetes Technology & Therapeutics8(3)2001

## Slide 41: Blood Volume Control (BVC)
Notes: دستگاه هایی که بیشتر از 2 الکترود در نوار تست قند خون داشته باشند معمولا مجهز به سیستم BVC هستند. یعنی اگر خون کم به دستگاه داده شود یا دست فرد بلرزد، errorمی دهد که در مورد گلوکوکارد صفر یک E7 گزارش می شود. ولی بسیاری از دستگاه هایی که هم اکنون به صورت رایگان عرضه می شود، این سیستم را ندارند و اگر خون کم به دستگاه داده شود به جای خطا، عدد نادرست گزارش می کنند. اینکه از کجا بفهمیم که دستگاه BVC دارد یا نه، باید بروشور دستگاه را بخوانیم. اگر حرفی از خطا با دادن خون کم نزده بود و به مصرف کننده گفته بود که خودتان کنترل کنید که محفظه نوار به میزان لازم از خون پر شده باشد، نشان دهنده این است که دستگاه BVC ندارد.

## Slide 42
Notes: در این مرحله اگر تعداد افراد شرکت کننده کم است، می توان یک نوار را بهشان نشان داد و بخواهیم که الکترودها را بشمارند.

## Slide 43
Notes: شاید برخی بروشور دستگاه ها مشخص نباشد که 2 الکترود دارد یا بیشتر. ولی راه دیگر برای درک این مطلب که دستگاه سیستم کنترل حجم خون دارد یا نه این است که در بروشور نوشته باشد که به ازای دادن خون ناکافی «خطا» می دهد نه عدد اشتباه. ولی دستگاه های نامعتبر از مصرف کننده می خواهد که خودش نوار را چک کند که اگر خون کامل محفظه را پر نکند، دوباره تست قند خون انجام دهد.

## Slide 44: Temperature
BG measurements with currently available test strips are temperature-dependent.
patients should be instructed to use BG systems within the specified operating temperature range only
Built-in temperature sensor
Erbach, M., Freckmann, G, et al (2016). Interferences and Limitations in Blood Glucose Self-Testing: An Overview of the Current Knowledge. J Diabetes Sci Technol, 10(5), 1161-1168.
Notes: Like every biochemical process, the test strip reaction during glucose measurement is influenced by temperature.34
Therefore, BG measurements with currently available test strips is temperature-dependent. BG performance under conditions that do not comply with the approved usage may result in erroneous BG measurements. For this reason, patients should be instructed to use BG systems within the specified operating temperature range only.22 Nowadays, many modern BG systems have a built-in temperature sensor
A study on 9 SMBG systems available in Norway explored the impact of temperature changes on the accuracy of BG measurements.
A change from 5°C to room temperature immediately before measurement, produced upward discrepancies >5% in 4 of these SMBG systems. Conversely, after a rapid change from 30°C to room temperature, 4 of the 9 BG systems presented downward discrepancies >5%.
Rapid extreme changes in temperature, however, may be associated with a time lag of limited function from 15 to 30 minutes.
After cold-stressing (−21°C) and heat-stressing (40°C) for up to 4 weeks, glucose test strips showed an impaired performance, with falsely elevated results after heating and falsely decreased results after cooling

## Slide 45: BGM measurement technology
Reflectance (photometric)
Electrochemical (amperometric)
William A Alto et al, Assuring the Accuracy of Home Glucose Monitoring, Diabetes Technology & Therapeutics 8(3) 2001
Notes: All disposable blood glucose meters function by one of two mechanisms: reflectance (photometric) or electrochemical (amperometric). Meters using reflectance measure a change in color on the strip that occurs after an enzymatic reaction with whole blood. Meters using electrochemical mechanisms work with small volumes of blood to measure electrochemical potential based on glucose oxidationاین روش با 2 تکنولوژی انجام می شود: فتومتری (یک ماده رنگی را اندازه گیری می کند) و یا الکتروشیمیایی (جریان الکتریکی را اندازه گیری می کند).
برای اینکه مفهوم فتومتری و الکتروشیمیایی خوب جا بیفتد، حتما یک دستگاه اکیوچک اکتیو هم داشته باشید. که مفهوم نور خوب جا بیفتد. با این عنوان، که این دستگاه در نسل خودش دستگاه خوبی است فقط حواسمون باشد که حداقل هر 2 ماه یکبار، درپوش آن را برداریم و روی عدسی آن را تمیز کنیم . در غیر اینصورت، اعداد خوانده شده از دستگاه، قابل قبول نیست.
ولی دستگاه ما نوری داشت، نه. پس ما مطمئن هستیم که بیمار ما از این جهت دچار خطا نمی شود.

## Slide 47: Thank you for your attention
Farir Asa Teb
Medical Affaires